## Slide 1
ABBA firmware and readout status
Alessandra Camplani – Università degli Studi di Milano
27/09/2016

## Slide 2
Outline
Firmware status
  New versions and their features
    General vision on each release produced
    New reset structure in front FPGAs
    New counters and registers for debugging
Settings status
  Latency settings
    Differences between different scenarios
  Number of samples study
    Campaing of calibration to clarify the loss of events saw in the past
Conclusion
2

## Slide 3
Firmware status
3

## Slide 4
Firmware version
Last presented version:
Front FPGA: version 0.96
Back FPGA: version 0.69a
For more details from the last LAr week: here
Current versions used for data taking:
Front FPGA: version 0.96e2
Back FPGA: version 0.69b
4

## Slide 5
Front FPGA version details
Version 0.96b
  Solved all the remaining timing constraints (thanks to Kenta!). Data HOP issue (spotted by Laura) solved.
Version 0.96c
  Improved reset structure: several soft_resets are now available to reset different interfaces and counters (more details in the next slides).
  New registers implemented.
Version 0.96d
  Merging and reworking some common vhdl files.
Version 0.96e/0.96e2
  Introduced new entity for ADC circular buffer.
Version 0.96f
  Substituted FIFOs and RAM with generic components (thanks to Philipp!). No changes in the firmware functionality, so not yet used in USA15.
5

## Slide 6
Back FPGA version details
Version 0.69a
  Added new pair of IP and MAC addresses for a new ABBA installed in USA15 (already presented in the last Lar week).
Version 0.69b
  Merged and reworked some common vhdl files.
Version 0.69c
  Substituted FIFOs and RAM with generic components (thanks to Philipp!). No changes in the firmware functionality, so not yet used in USA15.
6

## Slide 7
New reset structure in front FPGA
A new reset structure has been implemented and is now available in front FPGA firmware.
The signal in charge of this is the soft_resets signal (32 bit vector) where each bit is able to reset a specific register/counter/memory in front FPGA firmware (more detailed table in the next slide).
These resets can be used from the standalone tool, selecting the command “reset_front” and the FPGA:
./reset_front 10.145.91.19 1
A hexadecimal parameter will be required for the reset, for example:
1f to reset LTDB interface
7f00 to reset ABBA configuration
1000000 to reset the TTC decoder
60000000 to reset the interface to the back FPGA
depending on what has to be reset.
These resets are now also implemented in the ABBA package at CONFIG and START level.
7

## Slide 8
Available soft_resets bits
byte | bit | reset name | dedicated to |  | byte | bit | reset name | dedicated to
0 | 0 | rst_int_ltdb | LTDB interface |  | 2 | 16 |  | 
 | 1 | rst_mem_ltdb |  |  |  | 17 |  | 
 | 2 | rst_reg_ltdb_ctrl |  |  |  | 18 |  | 
 | 3 | rst_reg_ltdb_dec |  |  |  | 19 |  | 
 | 4 | rst_cnt_ltdb_err |  |  |  | 20 |  | 
 | 5 |  |  |  |  | 21 |  | 
 | 6 |  |  |  |  | 22 |  | 
 | 7 |  |  |  |  | 23 |  | 
1 | 8 | rst_reg_l1a_latency | ABBA configuration |  | 3 | 24 | rst_int_ttc | TTC interface
 | 9 | rst_mem_abba |  |  |  | 25 | rst_int_AV | avalon interface
 | 10 | rst_reg_header |  |  |  | 26 | rst_int_XAUI | XAUI interface
 | 11 | rst_cnt_l1a |  |  |  | 27 | rst_int_osc01 | (not yet implemented)
 | 12 | rst_cnt_pkg_write |  |  |  | 28 | rst_int_osc23 | oscillator interface
 | 13 | rst_cnt_pkg_diff |  |  |  | 29 |  | 
 | 14 | rst_cnt_pkg_sent |  |  |  | 30 |  | 
 | 15 |  |  |  |  | 31 |  | 
This is also
in the twiki under
the section firmware!
8

## Slide 9
New registers implemented
Two new registers have been implemented in the front FPGA firmware to monitor the packets:
Register name | Register address | Internal variable name | Read_counters interface name
L1A_TTYPE_CNT_REG_ADD* | 16011 | l1a_ttype_cnt | l1a received counter
L1A_ADC_PACKET_WR_CNT_REG_ADD | 16013 | l1a_adc_packet_wr_cnt | l1a packet built counter
L1A_ADC_PACKET_RD_CNT_REG_ADD | 16014 | l1a_adc_packet_rd_cnt | l1a packet trasmitted counter
* was called L1A_TTC_CNT_STATUS_REG_ADD and was used for both l1a received counter and l1a packet built counter.
Also in this case, the counters can be displayed from the standalone tool, with the command “read_L1A_counters”, selecting the FPGA:
./read_L1A_counters 10.145.91.19 1
They are also available now on the ABBA partition panel per each FPGA that is recording data.
9

## Slide 10
Setting status
10

## Slide 11
Latency and n°samples changes
Two major changes in the default settings for ABBA partition have been done:
On the 28th of June the number of samples read out changed to the actual value: 50 (instead of 60)
  Changed done to because we were losing 50% of data
  After this change data lost were only 5% (after 200 ... 300 events)
On the 15th of July the latency settings changed to the actual values (more details later):
  Changed done because the pulse tail was partially cut
  The aim was to have the pulse entirely in the new readout window (50 samples)
11

## Slide 12
Latency settings
Some calibration runs, with different latency settings, have been taken in USA15 for all four FPGAs.
For more details you can look here.
The position of the pulse (Lpeak) has been checked for each run.
The test confirmed that: a shift of 1 BC corresponds to a shift of 1 unit for the latency value.
So for example from an initial setting of 0x56 (86 in decimal):
to shift the pulse 10 BC to the right (in this window): increase the latency setting by 10, hence set it to 0x60 (96 in decimal)
12

## Slide 13
System latency differences
For the following different scenarios, the very same latency setting will show up the pulse in different positions:
Physics runs vs. calibration runs in USA15
  The difference between the pulse peak in physics runs and calibration runs in USA15 is of ~10 BC:
  LPhypeak + 10 BC = LCalibpeak
BNL LTDB vs. LAL LTDB in USA15
  The difference between the pulse peak from the BNL LTDB and from the LAL LTDB in USA15 is of ~8 BC:
  LBNLpeak + 8 BC = LLALpeak
Calibration in EMF vs. calibration in USA15
  The difference between the pulse peak in calibration runs in EMF and in USA15 is of ~15 BC:
  LEMF-Calibpeak + 15 BC = LUSA15-Calibpeak
Documented in the Twiki !!
13

## Slide 14
System latency now
BNL LTDB vs. LAL LTDB in USA15
  The difference between the pulse peak from the BNL LTDB and from the LAL LTDB in USA15 is of ~8 BC:
  LBNLpeak + 8 BC = LLALpeak
Latency settings have been changed to have all the pulses recorded with ABBA starting at around sample 20:
  The two different LTDBs now have specific settings (here all the details):
  - 0x4e for LAL LTDB (FPGAs 18:2 and 19:1)
  - 0x56 for BNL LTDB (FPGAs 19:2 and 20:2)
14

## Slide 15
Event loss during data taking
A campaign of calibration has been done to understand why there was a loss of events during the data taking.
For more details about the calibration, you can check here.
As a first attempt, the number of samples was decreased. All four FPGAs used.
The total number of events to be collected was 3000.
The L1A rate was at its minimum value: 40 Hz.
The influence of the number of samples on the recorded events percentage is not clearly visible.
BUT at the same time a strange behaviour of FPGA 19:1 was noticed.
 | Number of samples |  |  |  |  |  | 
 | 40 | 35 | 30 | 20 | 10 | 5 | 1
Event collected | 93% - 95% | 93% - 95% | 93% - 95% | 93% - 95% | 93% - 95% | 30% - 55% | 30% - 55%
15

## Slide 16
Calibration without FPGA 19:1
Inconsistency of the counters in FPGA 19:1 (same counters presented before)
L1a_received = 0
packet_built = 3000
packet_transmitted = 0
What is expect is the very same value everywhere if all the events are recorded.
If not all events are recorded, L1a_received number is bigger than the other two.
So, as a second attempt, new calibrations were taken, excluding FPGA: 19:1.
A new set of calibrations was collected and this time the remaining 3 FPGAs were all collecting the 100% of the events. The trigger rate was again 40 Hz.
The number of samples used were: 40, 50, 60 (in all the cases was reached the 100% of events recorded).
The conclusion of this test was that FPGA 19:1 is, for some reason, misbehaving and this seems to be the origin of the issue with the event recording.
16

## Slide 17
Calibration at different trigger rates
In the last set of calibration different trigger rates and number of samples were tested.
Only 3 FPGAs used: 19:1 excluded. Again the total number of events is 3000.
For more details about the calibrations, you can check here.
#samples | Trigger rate (Hz) |  |  |  |  |  |  |  | 
 | 80 | 120 | 160 | 200 | 240 | 280 | 320 | 360 | 400
60 |  |  | - | - | - | - | - | - | -
50 |  |  |  |  |  |  |  | - | -
40 |  |  |  |  |  |  |  | - | -
30 |  |  |  |  |  |  |  |  | 
20 |  |  | - | - | - | - | - | - | -
100% of events recorded
1 or 2 events lost per each FPGA
more than 50% events lost per each FPGA
"stop -> unconfig -> config -> start" used:
100% events collected
"stop -> unconfig -> config -> start" used:
1 or 2 events lost per each FPGA
17

## Slide 18
Conclusion
18

## Slide 19
Conclusion and outlook (1)
From the firmware point of view:
  Conclusion
  The data taking is stable, all four FPGAs are running
  The new features put in place allowed to spot new issues
  ToDo list:
  Work in progress for the UPD packet tracing between front and back FPGAs
  After the power cut and the installation of a new machine a new issue was spotted:
    If one of the bottom FPGAs (not involved in the data taking) are not programmed properly they will cause a (permanent) reset in the back FPGA (originating from the 'unused front' XAUI transceiver): as a consequence the board doesn't communicate.
    Done: programmed all the flash memory
    Todo: take an action on the reset structure of back FPGA
19

## Slide 20
Conclusion and outlook (2)
From the settings point of view:
  Latency settings:
  The pulses from both LTDBs are starting at the same sample now and the pulse is entirely in the window
  Number of samples:
  The campaign of calibrations shows that
    The influence of the number of samples on the recorded events percentage is not clearly visible
    The higher is the trigger rate the lower is number of events recorded
    Some strange behavior has been observed lowering the number sample (from 30 samples)
      To be investigated
20

## Slide 21
Thanks!
Any question?

## Slide 22
Backup slides

## Slide 23
Calibration at different trigger rates (part 1)
In the last set of calibration the trigger rate and the number of sample were changed.
For more details about the calibrations, you can check here.
In the first set of calibrations the settings were:
  50 samples
  frequency parameters set to 80, 120, 160, 200, 240, 280, 320 (a run per each value)
  Until 200: all the FPGAs were collecting 100% of events.
  For 240 and 280: 1 or 2 events lost per each FPGA.
  For 320: more than 50 events lost per each FPGA.
  (Just start and stop the ABBA partition per each run)
In the second set of calibrations the settings were:
  40 samples
  frequency parameters set to 80, 120, 160, 200, 240, 280, 320 (a run per each value)
  Until 200: all the FPGAs were collecting 100% of events.
  From 240 to 320: 1 or 2 events lost per each FPGA.
  (Just start and stop the ABBA partition per each run)
23

## Slide 24
Calibration at different trigger rates (part 2)
In the third set of calibrations the settings were
  30 samples
  frequency parameters set to 80, 120, 160, 200, 240, 280, 320, 360, 400 (a run per each value)
  This time the "start and stop" procedure for ABBA partition was not enough anymore.
  Every time that a new run was taken with just "start and stop" commands, 1 or 2 events got lost per each FPGA.
  So, instead of sequence of "stop -> start" I tried "stop -> unconfig -> config -> start" that contains a reset sent to all the FPGAs.
  With this procedure:
  Until 320: all the FPGAs were collecting 100% of events.
  For 360 and 400: 1 or 2 events lost per each FPGA.
In the fourth set of calibrations the settings were:
  - 20 samples
  - frequency parameters set to 80, 120 (a run per each value)
  This time was again different.
  Neither the procedure ("stop -> unconfig -> config -> start") allowed the FPGAs to collect all the data.
  NEVER reached the 100% of events collected.
24

## Slide 25
Calibration at different trigger rates (part 3)
In the fifth and last set of calibration the settings were:
  60 samples
  frequency parameters set to 80, 120 (a run per each value)
  For 80: all the FPGAs were collecting 100% of events ("stop -> start" procedure used).
  For 120: never reached the 100% of events collected (neither with "stop -> unconfig -> config -> start").
To summarize:
40 and 50 samples behaved as expected
20 and 30 samples not at all
60 sample maybe it’s simply a too high value
To be understood the reason of the strange behaviour for 20 and 30 samples.
25